## Slide 1: Principles of Neuronavigation: Frame and Frameless

## Slide 2: History
“Stereotactic”: From Greek “stereos”=3-dimensional and Latin “tactus”=to touch
1908: Horsely and Clarke develop first apparatus for insertion of probes into the brain based upon Cartesian planes and bony landmarks; only used in primates
1947: Spiegel and Wycis report first human use of stereotactic device . Goal was to perform minimally-invasive psychosurgery but first use was movement disorders
1948: Leksell develops first arc-centered frame
1957: Talairach publishes first atlas based upon ventriculography and intracranial brain landmarks rather than bone landmarks
1986: Kelley develops frame-based system for eye-tracking of operative microscope
1986: Roberts develops frameless acoustic-based system for tracking operative microscope
1991: Bucholz develops the first prototype for frameless sonic navigation of tracking tools and instruments in human cranial surgery
  Soon after incorporated optical digitizers to reduce inaccuracies from sound echoes

## Slide 3: General Principles of Stereotaxy
Navigation is based upon targeting relative to known reference points
Fiducial :
  From latin “fiducia” meaning trust
  A point of reference that can be visualized on imaging and identified by the surgeon and/or software package
  Accuracy of targeting is influenced by the number of fiducials around a target zone and the constancy of fiducials relative to the target
  Frame-based stereotaxy: Fiducials are bars built into cage or box that sits on frame during imaging
  Frameless stereotaxy: Fiducials are reference markers (stickers, bone screws) which are fixed directly to the patient prior to imaging

## Slide 4: Co-registration is the fundamental principle of stereotaxy
1906 -- Horsley & Clarke (animal) stereotactic frame

## Slide 5: Co-registration is the fundamental principle of stereotaxy
1906 – Horsley & Clarke (animal) stereotactic frame
1947 – Spiegel & Wycis (human) stereotactic frame

## Slide 6: Co-registration is the fundamental principle of stereotaxy
1906 – Horsley & Clarke (animal) stereotactic frame
1947 – Spiegel & Wycis (human) stereotactic frame
1947-1980 – Proliferation of stereotactic frames

## Slide 7: Co-registration is the fundamental principle of stereotaxy
1906 – Horsley & Clarke (animal) stereotactic frame
1947 – Spiegel & Wycis (human) stereotactic frame
1947-1980 – Proliferation of stereotactic frames
1980s – Computational resources enable “frameless” transformation-based stereotactic systems

## Slide 8: Considerations with Frame-Based Stereotaxy
Method of target localization
  Indirect vs. direct
Imaging errors due to frame placement
Imaging errors due to distortion

## Slide 9: Methods of Image-Based Target Localization
Indirect (Based upon position of AC-PC)
  Standard coordinates
    Leksell’s pallidotomy target is classic example
  Adjusted map
    Schaltenbrand-Wahren is most common
    Average AC-PC distance is 23-27mm; greater than 30mm should raise accuracy concerns
Direct (Target visually chosen from scan)

## Slide 10
AC
PC
Colliculi
AC-PC: Sagittal T2 Localizer

## Slide 11
PC
AC
Axial T2 Measurement of AC-PC

## Slide 12: Indirect Targeting: Fixed Coordinates
Thalamus (Vim)
  1-7mm posterior
  0-3mm superior
  12-17mm lateral
GPi
  2-3mm anterior
  3-6mm inferior
  18-22 mm lateral
STN
  3-5mm posterior
  5-6mm inferior
  11-14mm lateral
(All points relative to midcommissural point)

## Slide 13: Indirect targeting: Adjusted Map

## Slide 14: Direct targeting: STN

## Slide 15: Sources of Error: MRI Image Distortion
Magnetic field inhomogeneities and non-linear magnetic field gradients cause distortion
  Distortion often worst in coronal sections; measuring Leksell fiducials can determine distortion severity
Frame may introduce additional distortion
  Measuring target distance from MCP on preop MRI can guide targeting from framed image
CT not subject to these distortions; CT/MRI fusion may minimize effects of distortion
Bandwidth can influence contrast
  Lower bandwidth increases gray/white contrast to a point
  Very low bandwidth can worsen distortion

## Slide 16: Image Fusion

## Slide 17: Eight Things Every Neurosurgery Resident Should Know about Frameless Image-Guidance

## Slide 18: What is image-guided surgery and how does it work?
Image-guided surgery (neuronavigation, “frameless stereotaxy”) is an operative technique by which correlation between imaging studies and the operative field is provided.
This is accomplished by co-registration of imaging studies with the OR patient.

## Slide 19: Matrix Expression
P2I = 2ITM MTW WT3I P3I
Head in the world space, W
Microscope
frame, M
Image overlay, 2I
2ITM
Microscope calibration
MTW
Tracking patient and microscope
WT3I
Patient registration
Polaris
Preop. 3D image, 3I
Notes: The real power of computational stereotaxy lies in its ability to readily calculate the transformations between the various coordinate systems of imaging studies, operating rooms, microscopes, and just about anything else of interest.

## Slide 20: What equipment is involved?
Localization device (digitizer)
  e.g., optical, electromagnetic, articulated arm
  most systems today include a reference frame to enable OR table movement

## Slide 21: What equipment is involved?
Localization device (digitizer)
  e.g., optical, electromagnetic, articulated arm
Computer with registration algorithm

## Slide 22: What equipment is involved?
Localization device (digitizer)
  e.g., optical, electromagnetic, articulated arm
Computer with registration algorithm
Effector
  e.g., pointer and monitor, microscope heads-up display

## Slide 23: What types of co-registration strategies can be used?
Paired-point rigid transformation
Surface (contour) matching

## Slide 24: Fiducial registration error (FRE)
the root- mean square distance between corresponding fiducial points after registration
Fitzpatrick & West, 2001
Some important definitions…

## Slide 25: Fiducial localization error (FLE)
the error in locating the fiducial points
Fitzpatrick & West, 2001

## Slide 26: Target registration error (TRE)
the distance between corresponding points other than the fiducial points after registration
Fitzpatrick & West, 2001
This is what really matters!

## Slide 27: Accuracy in phantom testing
Benardete et al, 2001

## Slide 28: Clinical application accuracy (comparing seven registration methods)
Mascott et al, 2006

## Slide 29: What are the sources of error?
Imaging data set
  resolution
  e.g., slice thickness, pixel/voxel size
  spatial infidelity
  e.g., magnetic field inhomogenieties in echo planar fMRI
  imaging study fusion
  e.g., CT–MRI, atlas–MRI

## Slide 30: Dependence of stereotactic accuracy on image slice thickness
Maciunas et al, 1994
Notes: This study was done for stereotactic frames but the principle is the same,
Importance points:
No frame is submillimetric
Clinically important is C.I. (range)
Slice thickness plays a major role

## Slide 31: What are the sources of error?
Imaging data set
  resolution
  e.g., slice thickness, pixel/voxel size
  spatial infidelity
  e.g., magnetic field inhomogenieties in echo planar fMRI
Sumanaweera, 1994

## Slide 32: What are the sources of error?
Imaging data set
  resolution
  e.g., slice thickness, pixel/voxel size
  spatial infidelity
  e.g., magnetic field inhomogenieties in echo planar fMRI
  imaging study fusion
  e.g., CT–MRI, atlas–MRI

## Slide 33: What are the sources of error?
Imaging data set
Registration process (image–OR space)
  axes orientation (handedness of coordinate system)
  algorithm ambiguity
  fiducial number, configuration, displacement, OR localization (surgeon & digitizer)

## Slide 34: Number of fiducials and accuracy
Fitzpatrick et al, 1998
Steinmeier et al, 2000
Notes: Phantom studies, empirical observation, and simulation studies have all shown increasing accuracy with increasing number of fiducials.

## Slide 35
West et al, 2001
Notes: FIGURE 2. A–H, views of a large square and large triangular fiducial configurations and their corresponding TRE profiles. Missing iso-error contours indicate that the TRE value was either consistently above or below that isocontour value for the entire displayed slice. A, left and right view; B, coronal view; C, sagittal view; D, axial view; E, left and right view; F, coronal view; G, sagittal view; H, axial view.

## Slide 36: TRE has an approximate N-1/2 dependence
Fitzpatrick et al, 1998

## Slide 37: Error increases as the distance of the target from the fiducial centroid
West et al, 2001
Notes: As the distance of the target from the fiducial centroid increases, the error increases, approaching dependence. The error also increases as the configuration of the fiducials becomes smaller.

## Slide 38: FRE is not a reliable indicator of registration accuracy (!!)
FRE is independent of fiducial configuration
FRE is independent of bias errors (e.g., MRI gradient, digitizer camera malalignment, bent handheld probe)
Fitzpatrick et al, 1998

## Slide 39: Tips regarding fiducials
Avoid linear fiducial configurations
Arrange fiducials so that the center of their configuration is close to the region of interest during surgery
Spread out the fiducials
Use as many fiducials as reasonably possible
Mark scalp at fiducial site
Avoid occipital region or distorted scalp
partially adapted from West et al, 2001

## Slide 40: What are the sources of error?
Imaging data set
Registration process (image–OR space)
Digitizer performance
Wang & Song, 2011

## Slide 41: What are the sources of error?
Surgical field displacement or deformation
Dorward et al, 1998
Hill et al, 1998
Roberts et al, 1998
Ji et al, 2012
Notes: Numerous studies showing displacements, often exceeding 10mm.

## Slide 42: How does this relate to intraoperative MRI/CT?
Numerous implementations
Facilitated co-registration
Updated image data-set
Cost-benefit analyses pending

## Slide 43: In what applications has image-guidance been important?
Tumor (biopsy, resection of glial and met tumor)
Epilepsy (structural & physiologic data, resection)
Functional (DBS)
Spine (instrumentation)
Radiosurgery (frameless technologies)
Cerebrovascular (?)
Other: ENT, Plastics, Ortho, General

## Slide 44: What’s under development for image-guidance?
Automated registration
Ease of use
Updated imaging/registration
Increasing accuracy
Robotics
Extension of application to other
  surgeries, other disciplines
Nathoo, 2005
Louw, 2004